SUPPLEMENT 2: EVIDENCE TABLE FOR GLUCOSE METERS

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design |  | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) |  | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SHALL PORTABLE METERS BE USED IN DIAGNOSIS AND SCREENING OF DIABETES MELLITUS? |  |  |  |  |  |  |  |  |  |  |
| a | Portable glucose meters should not be used in the diagnosis of diabetes, including gestational diabetes. B (moderate) | There are insufficient published data to support a role for portable meters and skin-prick (finger-stick) blood samples in the diagnosis of diabetes or for population screening. C (moderate) The imprecision of the results, coupled with the substantial differences among meters, precludes the use of glucose meters in the diagnosis of diabetes and limits their usefulness in screening for diabetes. A (moderate) |  | American Diabetes Association. 2. Classification and Diagnosis of Diabetes: Standards of Medical Care in Diabetes-2022. Diabetes Care. 2021 Dec;45(Suppl 1):S17-S33. | Evidence-based guidance | Moderate |  | Moderate | WHO and USPSTF primarily recommend plasma, but WHO accepts capillary whole blood using glucose meter. WHO and ADA accept glucose meters for screening for practical and financial reasons in specific situations such as resource-poor areas and regions where access of patients to laboratory testing is impractical and where an adequate referral system is established for following up positive test results. This represents a strong consensus view that screening by glucometer is ”better than doing nothing”. Most glucose meters are not accurate enough to diagnose diabetes. This represents strong agreement of experts. Quality of evidence downgraded for inconsistency and indirectness of evidence on the use of glucose meters for screening/diagnosing DM. |  |
|  |  |  |  | Siu AL on behalf of U.S. Preventive Services Task Force. Abnormal Blood Glucose and Type 2 Diabetes Mellitus: Screening. Ann Intern Med. 2015;163:861-868 | Evidence-based guidelines | High |  |  |  |  |
|  |  |  |  | Siu AL. Screening for Abnormal Blood Glucose and Type 2 Diabetes Mellitus: U.S. Preventive Services Task Force Recommendation Statement. Annals of Internal Medicine. 2015 Dec 1;163(11):861–9. |  |  |  |  |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SHOULD PORTABLE METERS BE USED IN MONITORING TYPE 1 DIABETES MELLITUS? |  |  |  |  |  |  |  |  |
| b | Frequent blood glucose monitoring (BGM) is recommended for all people with diabetes who use intensive insulin regimens (with multiple daily injections or insulin pump therapy) and who are not using continuous glucose monitoring (CGM). A (high) | Self-monitoring of blood glucose (SMBG) is recommended for all insulin-treated patients with diabetes. A (high) | Refinement of population to whom the recommendation applies | DCCT Research Group. The effect of intensive treatment of diabetes on the development and progression of long-term complications in insulin-dependent diabetes mellitus. N Engl J Med 1993;329:977-986. | RCT | High | High | Intensive glycemic control in patients with type 1 diabetes was achieved in the DCCT by participants performing SMBG at least four times per day, hence the ADA recommendation and a strong consensus for SMBG to be performed four or more times per day in type 1 diabetes. |
|  |  |  |  | UK Prospective Diabetes Study (UKPDS) Group. Intensive blood-glucose control with sulphonylureas or insulin compared with conventional treatment and risk of complications in patients with type 2 diabetes (UKPDS 33). Lancet. 1998;352(9131):837–53. | RCT | High |  |  |
|  |  |  |  | Self-Monitoring of Blood Glucose. Diabetes Care. 1996;19(Supplement_1):S62–6. | Consensus statement | Low |  |  |
|  |  |  |  | Gerich JE, Mokan M, Veneman T, Korytkowski M, Mitrakou A. Hypoglycemia unawareness. Endocr Rev. 1991;12(4):356–71. | Review | Low |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SHOULD PORTABLE METERS BE USED IN MONITORING TYPE 2 DM? |  |  |  |  |  |  |  |  |
| c | Routine use of BGM is not recommended for people with type 2 diabetes treated with diet and/or oral agents alone. A (high) | In patients with type 2 diabetes treated with diet and oral agents, SMBG may help achieve better control, particularly when therapy is initiated or changed. Data are insufficient, however, to claim an associated improvement of health outcomes. The role of SMBG in patients with stable type 2 diabetes controlled by diet alone is not known. C (high) | New high level of evidence became available to refine recommendation | American Diabetes Association. 7. Diabetes Technology: Standards of Medical Care in Diabetes-2011. Diabetes Care. 2011;34(Suppl 1):S11-61. | Evidence-based guidance | Moderate | High | New meta-analyses, RCTs, evidence base guidance and a pragmatic RCT support the conclusion that the clinical benefit, if any, of routine use of SMBG in this patient group is small, transient, and accompanied by continuing costs. |
|  |  |  |  | Simon J, Gray A, Clarke P, Wade A, Neil A, Farmer A; Diabetes Glycaemic Education and Monitoring Trial Group. Cost effectiveness of self monitoring of blood glucose in patients with non-insulin treated type 2 diabetes: economic evaluation of data from the DiGEM trial. BMJ. 2008;336(7654):1177-80. | RCT and cost-effectiveness analysis | High |  |  |
|  |  |  |  | Malanda UL, Welschen LMC, Riphagen II, Dekker JM, Nijpels G, Bot SDM. Self-monitoring of blood glucose in patients with type 2 diabetes mellitus who are not using insulin. Cochrane Database Syst Rev 2012;1:CD005060 | Systematic review of RCTs | High |  |  |
|  |  |  |  | Farmer AJ, Perera R, Ward A, et al. Meta-analysis of individual patient data in randomised trials of self-monitoring of blood glucose in people with non-insulin treated type 2 diabetes. BMJ. 2012;344:e486. | Meta-analysis of individual patient data in RCTs | High |  |  |
|  |  |  |  | Machry RV, Rados DV, Gregório GR, Rodrigues TC. Self-monitoring blood glucose improves glycemic control in type 2 diabetes without intensive treatment: A systematic review and meta-analysis. Diabetes Res Clin Pract. 2018;142:173-187. | Meta-analysis of RCTs | High |  |  |
|  |  |  |  | Mannucci E, Antenore A, Giorgino F, Scavini M. Effects of structured vs unstructured self-monitoring of blood glucose on glucose control in patients with non-insulin-treated type 2 diabetes: a meta-analysis of randomized controlled trials. J Diabetes Sci Technol 2018;12:183–189 | Meta-analysis of RCTs | High |  |  |
|  |  |  |  | Young LA, Buse JB, Weaver MA, et al. Glucose Self-monitoring in Non–Insulin-Treated Patients With Type 2 Diabetes in Primary Care Settings: A Randomized Trial. JAMA Intern Med. 2017;177(7):920–929. | RCT | High |  |  |
|  |  |  |  | American Diabetes Association. 7. Diabetes Technology: Standards of Medical Care in Diabetes-2022. Diabetes Care. 2021 Dec;5(Suppl_1)11):S97-112. | Evidence-based guidance | Moderate |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design |  | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) |  | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| WHAT ARE THE PRE-ANALYTICAL CONSIDERATIONS FOR GLUCOSE METERS? |  |  |  |  |  |  |  |  |  |  |
| d | Individuals with diabetes should be instructed in the correct use of glucose meters, including technique of sample collection and use of quality control. GPP | Patients should be instructed in the correct use of glucose meters, including quality control. Comparison between SMBG and concurrent laboratory glucose analysis should be performed at regular intervals to evaluate the performance of the meters in the patient’s hands. B (moderate) |  | Ellison JM, Stegmann JM, Colner SL, Michael RH, Sharma MK, Ervin KR, Horwitz DL. Rapid changes in postprandial blood glucose produce concentration differences at finger, forearm, and thigh sampling sites. Diabetes Care. 2002 Jun;25(6):961-4. | Observational |  | High | Moderate |  | Downgraded recommendation to a good laboratory practice point for inconsistency and indirectness of evidence to health outcomes and because the recommendation addresses the analytical quality and organizational aspects of using glucose meters |
|  |  |  |  | Kabadi UM, et al. The effect of recurrent practice at home on the acceptability of capillary blood glucose readings. Accuracy of self blood glucose testing. Diab Care 1994;10:1110-23. | Observational |  | Moderate |  |  |  |
|  |  |  |  | American Diabetes Association. 7. Diabetes Technology: Standards of Medical Care in Diabetes-2022. Diabetes Care. 2021;45(Suppl 1):S97-112. | Guideline |  | Low |  |  |  |
| WHAT ARE THE ANALYTICAL CONSIDERATIONS FOR GLUCOSE METERS? |  |  |  |  |  |  |  |  |  |  |
| e | Glucose meters should report the glucose concentrations in plasma rather than in whole blood to facilitate comparison with plasma results of assays performed in accredited laboratories. GPP | Meters should measure and report plasma glucose concentrations to facilitate comparison with assays performed in accredited laboratories. GPP | Refined text for clarity | Burnett RW, D’Orazio P, Fogh-Andersen N, Kuwa K, Kulpmann WR, Larsson L, et al. IFCC recommendation on reporting results for blood glucose. Clin Chim Acta 2001;307:205–9 | Expert opinion | Low |  | Low | Downgraded recommendation to a good laboratory practice point for low quality and indirectness of evidence to health outcomes and because the recommendation addresses the technical and organizational aspects of point of care testing using glucose meters. |  |
|  |  |  |  | D’Orazio P, Burnett RW, Fogh-Andersen N, Jacobs E, Kuwa K, Kulpmann WR, et al. Approved IFCC recommendation on reporting results for blood glucose (abbreviated). Clin Chem 2005;51:1573–6 | Expert opinion | Low |  |  |  |  |
|  |  |  |  | Steffes MW, Sacks DB. Measurement of circulating glucose concentrations: The time is now for consistency among methods and types of samples. Clin Chem 2005;51:1569–70 | Opinion | Low |  |  |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| f | Glucose meters should meet relevant accuracy standards of the FDA in the U.S. or comparable analytical performance specifications in other locations. GPP | Multiple performance goals for portable glucose meters have been proposed. These targets vary widely and are highly controversial. Manufacturers should work to improve the imprecision of current meters, with an intermediate goal of limiting total error, for 95% of samples, to ≤15% at glucose concentrations ≥5.6 mmol/L (100 mg/dL) and to <0.8 mmol/L (15 mg/dL) at glucose concentrations <5.6 mmol/L (100 mg/dL). Lower total error would be desirable and may prove necessary in tight glucose control protocols and for avoiding hypoglycemia in all settings. C (low) Studies are needed to determine the analytical goals (quality specifications) for glucose meters in self-monitoring of blood glucose and in intensive care units. C (moderate) | Simplified recommendations for clarity in view of recent evidence | U.S. Food and Drug Administration. Self-Monitoring Blood Glucose Test Systems for Over-the-Counter Use: Guidance for Industry and Food and Drug Administration Staff. 2020. Accessed 12 May 2021. https://www.fda.gov/media/87721/download | Regulatory Guidance | Low/ moderate | Low- moderate | Performance goal targets vary widely and are highly controversial. However, there is now increased level of agreement of experts in regulatory (FDA) and quality standards organizations (ISO and CLSI) that laboratories follow locally. Downgraded recommendation to a good laboratory practice point for inconsistency and indirectness of evidence to health outcomes and because the recommendation addresses the analytical aspects of using glucose meters |
|  |  |  |  | Blood Glucose Monitoring Test Systems for Prescription Point-of-Care Use [Internet]. U.S. Food & Drug Administration Center for Devices and Radiological Health; 2020 [cited 2022 May 23]. Available from: https://www.fda.gov/regulatory-information/search-fda-guidance-in vit documents/blood-glucose-monitoring-test-systems-prescription-point-care-use | Regulatory Guidance | Low/ moderate |  |  |
|  |  |  |  | International Standards Organization. ISO15197:2013. In vitro diagnostic test systems –requirements for blood glucose monitoring systems for self-testing in managing diabetes mellitus. Reviewed 2018. Accessed 12 May 2021. https://www.iso.org/standard/54976.html | Consensus Standard | Low/ moderate |  |  |
|  |  |  |  | POCT12-A3. Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facilities, 3rd Edition. 2013. CLSI. Annapolis Junction, MD. | Consensus Standard | Low/ moderate |  |  |
|  |  |  |  | Ekhlaspour L, Mondesir D, Lautsch N, Balliro C, Hillard M, Magyar K, Radocchia LG, Esmaeili A, Sinha M, Russell SJ. Comparative Accuracy of 17 Point-of-Care Glucose Meters. J Diabetes Sci Technol. 2017;11(3):558-566 | Analytical performance study | Moderate |  |  |
|  |  |  |  | Boyd JC, Bruns DE. Quality specifications for glucose meters: assessment by simulation modeling of errors in insulin dose. Clin Chem 2001; 47:209–14. | Simulation modeling | Low/ moderate |  |  |
|  |  |  |  | Campos-Náñez E, Fortwaengler K, Breton MD. Clinical Impact of Blood Glucose Monitoring Accuracy: An In-Silico Study. J Diabetes Sci Technol. 2017;11(6):1187–1195. | Simulation (in silico) study | Low-moderate |  |  |
|  |  |  |  | Fortwaengler K, Campos-Náñez E, Parkin CG, Breton MD. The Financial Impact of Inaccurate Blood Glucose Monitoring Systems. J Diabetes Sci Technol. 2018;12(2):318–324. | Simulation (in silico) study | Low-moderate |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ARE GLUCOSE METERS ADEQUATE FOR WIDESPREAD USE IN ACUTE CARE SETTINGS? |  |  |  |  |  |  |  |  |
| g | In hospitals and acute-care facilities, point-of-care testing personnel, including nurses, should use glucose meters that are intended for professional use. GPP | Recommendations for future research: Important end-points in studies of SMBG should include, at a minimum, HbA1c and frequency of hypoglycemic episodes to ascertain whether improved meters allow patients to achieve better glucose control. For studies of use in intensive or critical care, important end-points include mean blood glucose, frequency of hypogylcemia and variability of glucose control. Ideally, outcomes (e.g., long-term complications) should also be examined. GPP | 1. New meters are available and cleared for professional use. 2. New guidance issued by U.S. FDA | Nichols JH, Brandler ES, Fantz CR, Fisher K, Goodman MD, Headden G, Hoppensteadt D, Matika R, Peacock WF, Rodrigo J, Schützenmeister A, Swanson JR, Canada-Vilalta C, Miles G, Tran N. A Multicenter Evaluation of a Point-of-Care Blood Glucose Meter System in Critically Ill Patients. J Appl Lab Med. 2021;6(4):820-833. | Observational study | Low | Low | Downgraded recommendations to good laboratory practice points for indirectness of evidence to health outcomes and because the recommendation addresses the organizational aspects of care. |
|  |  |  |  | Mitsios JV, Ashby LA, Haverstick DM, Bruns DE, Scott MG. Analytic evaluation of a new glucose meter system in 15 different critical care settings. J Diabetes Sci Technol 2013;7:1282-7. | Analytical performance study | Low |  |  |
|  |  |  |  | Karon BS, Meeusen JW, Bryant SC. Impact of Glucose Meter Error on Glycemic Variability and Time in Target Range During Glycemic Control After Cardiovascular Surgery. J Diabetes Sci Technol. 2015;10(2):336–342. | Observational study with historical controls | Low |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| h | When testing newborns, personnel should use only meters that are intended for use in newborns. GPP | N/A | New recommendation for neonatal testing of glucose with POC devices | Warner JV, Wu JY, Buckingham N, McLeod DS, Mottram B, Carter AC. Can one point-of-care glucose meter be used for all pediatric and adult hospital patients? Evaluation of three meters, including recently modified test strips. Diabetes Technol Ther. 2011;13(1):55-62. | Analytical performance study | Low | Low | Glucose meters in neonatal care are preferred options due to small sample volume requirements. However, testing in a neonatal matrix poses special challenges to laboratories due to numerous interferences that impact on the accuracy of measurements. Only devices assessed for common interferences in neonatal specimens and proven to provide reliable results can be safely used in newborns. The evidence between analytical performance of these POC devices and patient management or outcomes is indirect, therefore analytical performance data provides a weak link to patient outcomes, which resulted in the downgrading of the body of evidence behind a good laboratory practice recommendation. |
|  |  |  |  | Ceriotti F, Kaczmarek E, Guerra E, Mastrantonio F, Lucarelli F, Valgimigli F, Mosca A. Comparative performance assessment of point-of-care testing devices for measuring glucose and ketones at the patient bedside. J Diabetes Sci Technol. 2015;9(2):268-77. | Analytical performance study | Low |  |  |
|  |  |  |  | Inman M, Kyle BD, Lyon ME. Contribution of Glucose Meter Error to Misclassification of Neonatal Glycemic Status. JAMA Pediatr. 2021;175(5):453-455. | Simulation modeling | Low |  |  |
|  |  |  |  | Wada Y, Nakamura T, Kaneshige M, Takahashi S, Fujinaga H, Tsukamoto K, et al. Evaluation of two glucose meters and interference corrections for screening neonatal hypoglycemia. Pediatr Int. 2015;57(4):603–7. | Analytical interference study | Low |  |  |
|  |  |  |  | Demircik F, Ramljak S, Hermanns I, Pfützner A, Pfützner A. Evaluation of hematocrit interference with MyStar extra and seven competitive devices. J Diabetes Sci Technol. 2015;9(2):262–7. | Analytical interference study | Low |  |  |
|  |  |  |  | Teodorczyk M, Cardosi M, Setford S. Hematocrit compensation in electrochemical blood glucose monitoring systems. J Diabetes Sci Technol. 2012;6(3):648–55. | Analytical interference study | Low |  |  |
|  |  |  |  | Tendl KA, Christoph J, Bohn A, Herkner KR, Pollak A, Prusa AR. Two site evaluation of the performance of a new generation point-of-care glucose meter for use in a neonatal intensive care unit. Clin Chem Lab Med. 2013;51(9):1747–54. | Analytical performance study | Low |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) |  | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HOW FREQUENTLY SHOULD PATIENTS WITH TYPE 1 DIABETES PERFORM SMBG? |  |  |  |  |  |  |  |  |  |
| i | Unless CGM is used, people using multiple daily injections of insulin should be encouraged to perform BGM at a frequency appropriate for their insulin dosage regimen, typically at least 4 times a day. B (moderate) | SMBG should be performed at least 3 times per day in patients with type 1 diabetes (ungraded position without recommendation) | Improve clarity and move from position statement to recommendation | American Diabetes Association Professional Practice Committee. 7. Diabetes Technology: Standards of Medical Care in Diabetes—2022. Diabetes Care. 2021 Dec 16;45(Supplement_1):S97–112. | Guideline expert opinion |  | Moderate | Moderate | There is moderate quality evidence and level of agreement of experts that the recommended frequency of monitoring improves health outcomes. Recommendation is downgraded to B although there is strong expert consensus regarding frequency of SMBG in this subgroup of patients. However, the outcome measure was improvement of HbA1c, which has an indirect link to health outcomes. |
|  |  |  |  | Miller KM, Beck RW, Bergenstal RM, et al.;T1D Exchange Clinic Network. Evidence of a strong association between frequency of self-monitoring of blood glucose and hemoglobin A1c levels in T1D exchange clinic registry participants. Diabetes Care 2013;36:2009–2014 | Observational study (Retrospective, cross-sectional) |  | Moderate |  |  |
|  |  |  |  | Ziegler R, Heidtmann B, Hilgard D, Hofer S,Rosenbauer J, Holl R; DPV-Wiss-Initiative. Frequency of SMBG correlates with HbA1c and acute complications in children and adolescents with type 1 diabetes. Pediatr Diabetes 2011;12:11–17 | Observational study (Retrospective, cross-sectional) |  | Low |  |  |
|  |  |  |  | Schütt M, Kern W, Krause U, Busch P, Dapp A, Grziwotz R, Mayer I, Rosenbauer J, Wagner C, Zimmermann A, Kerner W, Holl RW; DPV Initiative. Is the frequency of self-monitoring of blood glucose related to long-term metabolic control? Multicenter analysis including 24,500 patients from 191 centers in Germany and Austria. Exp Clin Endocrinol Diabetes. 2006 Jul;114(7):384-8. | Observational study (Retrospective, cross-sectional) |  | Low |  |  |

| No | 1. Updated/new recommendation with its grade and quality of evidence(2) | 2. NACB 2011 recommendation and its grade(1) | 3. Why was it necessary to modify the recommendation? | 4. Key references supporting the new recommendation | 5. Study design | 6. Level of evidence(2) (high-moderate-low) | 7. Quality of evidence(2) (high-moderate-low-very low) | 8. Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| j | Manufacturers should continue to improve the analytical performance of meters. GPP | N/A | New recommendation for manufacturers | Ekhlaspour L, Mondesir D, Lautsch N, Balliro C, Hillard M, Magyar K, Radocchia LG, Esmaeili A, Sinha M, Russell SJ. Comparative Accuracy of 17 Point-of-Care Glucose Meters. J Diabetes Sci Technol. 2017;11(3):558-566. | Analytical performance study | Moderate | Low |  |
|  |  |  |  | Yu-Fei W, Wei-Ping J, Ming-Hsun W, Miao-O C, Ming-Chang H, Chi-Pin W, Ming-Shih L. Accuracy Evaluation of 19 Blood Glucose Monitoring Systems Manufactured in the Asia-Pacific Region: A Multicenter Study. J Diabetes Sci Technol. 201711(5):953-965. | Analytical performance study | Moderate |  |  |
|  |  |  |  | McQueen RB, Breton MD, Ott M, Koa H, Beamer B, Campbell JD. Economic Value of Improved Accuracy for Self-Monitoring of Blood Glucose Devices for Type 1 Diabetes in Canada. J Diabetes Sci Technol. 2015 Aug 14;10(2):366-77. | Simulation study | Low-Moderate |  |  |
|  |  |  |  | McQueen RB, Breton MD, Craig J, Holmes H, Whittington MD, Ott MA, Campbell JD. Economic Value of Improved Accuracy for Self-Monitoring of Blood Glucose Devices for Type 1 and Type 2 Diabetes in England. J Diabetes Sci Technol. 2018;12(5):992-1001. | Simulation study | Low-Moderate |  |  |